## Slide 1: Interdigitated Electrodes in Microfluidics Erin Ingram & Omar Mahmood Dr. Arum Han, Dr. Han Zhang NanoBio Systems Lab
Notes: Erin

## Slide 2: Problem Statement
Satellite droplets form during droplet generation
Negatively affect the manipulation of desired droplets
  Cross contamination
  False test results
Notes: Omar

## Slide 3: Our solution
Design a two stage filtration system incorporating an Interdigitated Electrode (IDE) arrays to perform size based filtration and prevent satellite droplets from making it to the chip’s output
Fabricate entire device in AggieFab – nano/micro fabrication facility
Fluid flow
Electrodes
Notes: Erin

## Slide 4: Subsystem Updates

## Slide 5: Overall design
+
=
Notes: Erin

## Slide 6: Subsystems and Interface
Droplet
Channel - Fabricated by Omar
IDE- Fabricated by Omar
Satellite droplets - “waste”
Desired droplets
Various droplet sizes
IDE - Fabricated by Erin
1
2
3
1- Droplets not big enough - “Waste”
2- Droplets too big - “Waste”
3- Perfect size
Final Droplets
Satellite droplet
PCR
Channel - Fabricated by Erin
Notes: Both

## Slide 7: High Pass Filter - Omar
Progress:
3D printed master mold using Nanoscribe
  Troubleshot printing error
Fabricated channel copies via softlithography
Validating Stage 1
Notes: Omar

## Slide 8: Stage 1 Validation
Generated 85 µm droplets using a specialized generator
Droplets passed through channel successfully
  Testing print quality
  No leakages or clogs
Next step: validating Omar’s droplet generator
Notes: Omar

## Slide 9: Band Pass Filter - Erin
Progress:
Printing master mold – Nanoscribe under repair
COMSOL - electric field simulations
Plasma bonding
  Troubleshooting
    Out of oxygen
  Trained Omar
Microscope training
Next steps: print master mold, softlithography, validate channel quality
Notes: Erin

## Slide 10: Execution Plan
 | 9/16/22 | 9/30/22 | 10/14/22 | 10/28/22 | 11/11/22 | 11/25/22
High Pass Filter (Omar) | Train on non-cleanroom related tasks: -Soft-lithography -Plasma cleaner -Microscopes -PCR techniques | Design the channel, fabricate master mold using NanoScribe | Design IDEs Fabricate softlithography copies of channel | Finish droplet generation validation of Stage 1 Plasma bonding Train on metal deposition, photolithography, wet etching | COMSOL simulations complete Perform experiments using the bonded device, experiment with different IDE characteristics | Validate high pass filtration using ideal IDE parameters
Band Pass Filter (Erin) | Train on non-cleanroom related tasks: -Soft-lithography -Plasma cleaner -Microscopes -PCR techniques | Design the channel, fabricate master mold using NanoScribe (UPDATE: on hold) | Design IDEs Plasma bonding training COMSOL training | Microscope/high speed camera training Train on metal deposition, photolithography, wet etching Plasma bonding Expecting NanoScribe to be fixed, master mold to be printed | COMSOL simulations complete Fabricate softlithography copies of channel Perform experiments using the bonded device, experiment with different IDE characteristics | Validate band pass filtration using ideal IDE parameters
Notes: Both

## Slide 11: Validation Plan
Stage 1 Channel | Stage 2 Channel | High Pass IDE Filter | Band Pass IDE Filter | Finished Device
Print quality Droplet generation | Print quality Droplet generation | Verify electric field strengths Filtration of small droplets | Verify electric field strengths Filtration of small and large droplets | Successful droplet generation and filtration of all droplet sizes in both stages “Successful” = 98% accurate filtration
***Both team members follow the same schedule for their specific subsystem
Notes: Erin

## Slide 12: Thanks! Questions?